## Slide 1
New concepts and instruments for 14C measurements in i-graphite
EDF/DP2D - Gérard Laurent; IN Solutions – Jean Luc Dumont, Christophe Fongarland, Philippe Le Tourneur ; Airbus Defense and Space – Christophe Mathonat, Clément Deyglun, Alain Godot, Guillaume Jossens; Kep Nuclear

## Slide 2: New concepts and instruments for 14C measurements in i-graphite
Intro : The need of field measurement capabilities for the characterization of the i-graphite
What can be done for 14C measurements ?
  Calorimetry : new means
  Gamma response
  Estimated performances
Conclusion : next steps
Outlook
Sleeve (graphite)
Wires and Support (stainless steel)
Fuel
www.andra.fr

## Slide 3: New concepts and instruments for 14C measurements in i-graphite
14C (and 36Cl) are key nuclides for irradiated graphite waste management
  Key for designing methods, tools and decommissioning means
  Key for treatment and packaging
  Key for means and ways of storage…
Unfortunately they are difficult to measure. Need of sampling and laboratory
Complementary to lab measurements, field measurement means would and could be very useful : no sampling, real time measurement
Our goal, in this first step, is to propose and assess technical means apt to make this job
Next steps will be then first to confirm experimentally on sites then, if relevant, to early design and evaluate how future processes would take benefit of such kind of instruments
Intro

## Slide 4: New concepts and instruments for 14C measurements in i-graphite
Half life : 5730 years
Mostly created by neutron capture of stable isotope 13C or (n,p) reaction on 14N
Pure beta emitter (Emean 50 keV, 100%)
Level of activity in i-graphite : about 104 to 105 Bq/g typical
We have decided to explore the conventional way of calorimetry but with a new instrument able to lower the usual detection limits
2 a priori drawbacks :
  Being able to measure very low heating powers
  Being able to subtract the other isotopes (activation products) heating contribution
14C

## Slide 5: New concepts and instruments for 14C measurements in i-graphite
µLVC Large Volume Calorimeter
KEP Nuclear developed a new Large Volume Calorimeter: µLVC
  Differential isothermal heat flow calorimeter
  Sample inside a vacuum chamber
  Effective power range 50 µW – 100 mW
  Measurement Mean Random Uncertainty on the first prototype:
40 L
Height 660 mm – Diam. 280 mm
Thermal bath
Measurement Chamber
Reference Chamber
Weight:
9 Tons
2600 mm
2400 mm
3000 mm
14C
Expected Uncertainties | Measured Uncertainties
±10% between 10 µW and 100 µW | ±20% between 50 µW and 250 µW
±1% between 100 µW and 1 mW | ±5% between 250 µW and 1 mW
±0.20% between 1 mW and 10 mW | ±1% between 1 mW and 100 mW

## Slide 6: New concepts and instruments for 14C measurements in i-graphite
Calorimetric assay of graphite is appropriate but:
  Relative low energy quantity released by 14C decay:
  50 keV in average  8E-15 W/Bq  8 µW per GBq.
  Irradiated graphite does not contain only 14C → obliged to calculate the contribution of other isotopes for correction
The identification of beta emitters is done by gamma spectrometry:
  Identifying charged particle emitters and quantifying gamma rays.
  Estimation of the power brought to the object by the gammas thanks to modeling.
14C

## Slide 7: New concepts and instruments for 14C measurements in i-graphite
The gamma-heating Issue
The energy deposit in graphite and calorimeter due to gamma photons is estimated using MCNPX code.
Measurement chamber
Thermal bloc
3D view of the µLVC
Reference chamber
Cross section of the µLVC
(MCNP Model)
Vacuum chamber
Intermediate bloc
2 MeV gamma rays source at 1E9 Bq (the worst case):
  The gamma-heating of the intermediate bloc = 20 µW
  The gamma-heating of the thermal bloc = 50 µW.
  Heat produced by the temperature sensors ~1 mW >> 50 µW
  → The heating of the regulation blocs is not measured during the assay, and does not disturb the calorimeter working.
14C

## Slide 8: New concepts and instruments for 14C measurements in i-graphite
The gamma-heating inside the graphite sleeve and the measurement chamber depends of the source location :
  In graphite sleeve: if the source is closed to the outside → gamma rays can easily leak → not induce heating inside the sleeve.
  The measurement chamber:
    If the source in close to the lateral wall → gamma rays go through the thin copper wall → not induce heating inside the chamber.
    If the source is close to the low part → gamma rays are absorbed by the thick back part of the chamber. → induce heating inside the chamber.
Gamma-Heating calculation
Chart: Gamma heating of the graphite sleeve
| 0.015 | 0.05 | 0.1 | 0.2 | 0.511 | 0.75 | 1.0 | 1.2 | 1.5 | 2.0
Position 1. | 2.2429e-09 | 2.3866e-09 | 3.4573e-09 | 7.1873e-09 | 1.8004e-08 | 2.4991e-08 | 3.144e-08 | 3.607e-08 | 4.241e-08 | 5.1837e-08
Position 2. | 1.3168e-09 | 1.2513e-09 | 1.914e-09 | 4.3022e-09 | 1.1223e-08 | 1.5551e-08 | 1.9486e-08 | 2.2291e-08 | 2.6136e-08 | 3.1884e-08
Position 3. | 2.3345e-09 | 2.6828e-09 | 4.073e-09 | 8.8695e-09 | 2.3929e-08 | 3.4254e-08 | 4.4062e-08 | 5.1251e-08 | 6.1221e-08 | 7.6224e-08
Position 4. | 2.2003e-09 | 1.6796e-09 | 2.6546e-09 | 6.1352e-09 | 1.7103e-08 | 2.452e-08 | 3.1501e-08 | 3.6586e-08 | 4.3618e-08 | 5.4181e-08
Uniforme distribution | 2.1999e-09 | 2.3451e-09 | 3.5938e-09 | 7.9274e-09 | 2.1535e-08 | 3.0848e-08 | 3.9682e-08 | 4.6149e-08 | 5.5111e-08 | 6.8586e-08
Chart: Gamma heating of the measurement chamber
| 0.015 | 0.05 | 0.1 | 0.2 | 0.511 | 0.75 | 1.0 | 1.2 | 1.5 | 2.0
Position 1. | 1.1091e-10 | 5.6124e-09 | 1.1365e-08 | 1.3576e-08 | 1.8961e-08 | 2.3909e-08 | 2.9001e-08 | 3.2902e-08 | 3.8601e-08 | 4.8072e-08
Position 2. | 1.0459e-09 | 6.7523e-09 | 1.3712e-08 | 2.2284e-08 | 4.4777e-08 | 6.1919e-08 | 7.9362e-08 | 9.285e-08 | 1.1249e-07 | 1.4447e-07
Position 3. | 6.5578e-11 | 5.3218e-09 | 1.1145e-08 | 1.3612e-08 | 1.8202e-08 | 2.2672e-08 | 2.7415e-08 | 3.1126e-08 | 3.6629e-08 | 4.5908e-08
Position 4. | 1.9999e-10 | 6.3274e-09 | 1.3169e-08 | 2.1551e-08 | 4.0838e-08 | 5.5206e-08 | 6.9818e-08 | 8.1138e-08 | 9.771e-08 | 1.2497e-07
Uniforme distribution | 1.961e-10 | 5.6562e-09 | 1.1557e-08 | 1.5001e-08 | 2.2395e-08 | 2.8716e-08 | 3.5257e-08 | 4.0327e-08 | 4.7775e-08 | 6.0168e-08
14C

## Slide 9: New concepts and instruments for 14C measurements in i-graphite
14C measurement estimated performances
Calorimeter characteristics
Accuracy :
50µW-250 µW : 20%
250 µW- 1mW : 5%
1 mW – 100 mW: 1%
14C
Existing proto
LD = 50 µW
Ideal Perf
14C real activity (Bq/g)
Graphite sleeve 20 kgs
Measured activity increased of its error
(Bq/g)
(typical value 105 Bq/g)
Chart: 
| 1000.0 | 100.0 | 10000.0 | 50000.0 | 100000.0 | 500000.0 | 1.0E6 | 5.0E6 | 1.0E7 | 5.0E7 | 1.0E8
2,00E+07 2,00E+06 2,00E+08 1,00E+09 2,00E+09 1,00E+10 2,00E+10 1,00E+11 2,00E+11 1,00E+12 2,00E+12 | 313490 | 312599 | 322400 | 362000 | 411500 | 594000 | 1188000 | 5197500 | 9999000 | 49995000 | 99990000
 | 1000 | 100 | 10000 | 50000 | 100000 | 500000 | 1000000 | 5000000 | 10000000 | 50000000 | 100000000
 | 125990 | 125099 | 134900 | 174500 | 224000 | 594000 | 1188000 | 5197500 | 9999000 | 49995000 | 99990000
 | 63490 | 62599 | 72400 | 112000 | 161500 | 594000 | 1188000 | 5197500 | 9999000 | 49995000 | 99990000
LD 50 µW
LD 20 µW
LD 10 µW
LD = 20 µW
LD = 10 µW
Gamma heating error not taken into account
(108 Bq/sleeve makes 100 µW, ie error about 10µW)
Next steps

## Slide 10: New concepts and instruments for 14C measurements in i-graphite
Conclusion & perspectives
14C
Feasibility of instruments dedicated to field measurements of 14C in rough (not any preparation or sampling) i-graphite objects has been assessed. On site real measurement tests remain to do.
Limits of detection yet to be worked for 14C through improvement of a new low heating power calorimeter.
Processes and waste management scenarios which could use such equipments have to be thought for evaluation of interest.

## Slide 11: New concepts and instruments for 14C and 36Cl measurements in i-graphite
Conventional positron measurement by a gamma spectro
36Cl
11
Gamma detector in front of the object.
Counting the 511 keV events to know the number of positons
Exploiting the spectrum to know present gamma emitters
Deducing from all that the 36Cl activity
n
electron pulse
scintillator light pulse
511 keV photon
Location where the positron becomes 2 511 keV gammas
36Cl location (where the b+ or positron is emitted)
n
PMT
scintilator
Cl36 (b+)
Graphite
sleeve
In short : - this would work well if 36Cl is the only isotope in the graphite
- It would work less good if
    36Cl is a marginal b+ producer (among others)
    High energy gamma flux is high (specific enemy: e+e- pair production)